Hook up Acquity (UPLC) or HPLC to Quattromicro

Version 1.0 by Junsung

Disconnect instrument

- HPLC from Quattromicro

  - Unconfigure system

    - Go to MassLynx

    - Inlet method → tools → instrument configuration → configure → next → check to “None” (pump, autosampler, detector) → next → finish

    - Turn off HPLC

    - Close MassLynx and shutdown computer

  - Disconnect lines

    - Disconnect cable of contact clouser in event 1 of Quattromicro

    - Disconnect GPIB cable of HPLC

    - Disconnect Ethernet cable in “PC Link” of Quattromicro

- UPLC from Quattromicro

  - Unconfigure system

    - Go to MassLynx

    - Inlet method → tools → instrument configuration → configure → next → check to “None” (pump, autosampler, detector) → next → finish

    - Turn off HPLC

    - Close MassLynx and shutdown computer

  - Disconnect lines

    - Disconnect GPIB cable (if don’t use PDA)

Connect instrument

1. HPLC + Quattromicro

- Connect lines

  - Connect clouser in event 1 cable from HPLC

  - GPIB cable connect to HPLC from computer

  - Connect Ethernet cable to Quattromicro from computer

(Ethernet cable port is signed “PC Link”)

- HPLC connect to power (don’t turn it on)

- Computer boot on

- Windows “start” → go to “run”

  - type “telnet epc”: type “reboot”. If it is ok, it will be a beep signal.

  - type “ping epc –t: compunication request time out time< 1ms TTL=64

  - If that same message would be filled in screen, it is Ok.

- Configuration system

  - Go to MassLynx

    - Inlet method → tools → instrument configuration → configure → next → check to “Waters 2695” of select Pump → next → check to “Waters 2695” of select Auto Sampler → next → check to “None” of select Detector (if don’t use PDA) → next → next → Finish

    - Go to Events & Triggering → next → check to “Event In 1” → next → Finish

    - Press Finish button of Inlet Configuration

  - Go to LC in lnlet Method → press “Reset Communicaitons”

- Connect eluent line

2. UPLC + Quattromicro

- Connect lines

  - Connect Ethernet cable to #2 switch box from “PC Link” of Quattromicro

  - Connect Ethernet cable to #1 switch box from PC

  - Connect start cable between ACQURITY (UPLC) and the PC.

- UPLC connect to power (don’t turn on)

- Computer Boot on

- Windows “start” → go to “run”

  - Type “telnet epc”: type “reboot”. If it is ok, it will be a beep signal.

  - Type “ping epc –t: compunication request time out time< 1ms TTL=64

  - If that message would be filled in screen, it is Ok.

- Configuration system

- Go to MassLynx

    - Inlet method → tools → instrument configuration → configure → next → check to “Waters Acquity” of select Pump → next → check to “Waters Actuity” of select Auto Sampler → next → check to “None” of select Detector (if don’t use PDA) → next → next → Finish

    - Go to Events & Triggering → next → check to “Event In 1” → next → Finish

    - Press Finish button of Inlet Configuration

- Go to LC in lnlet Method → press “Reset Communicaitons”

- Connect eluent line