Hook up Acquity (UPLC) or HPLC to Quattromicro

Version 2.0 by Junsung

Connect HPLC to Quattromicro

- Power off all machines

- Turn off UPLC, HPLC and computers except Quattromicro

- Connect the HPLC cable

- Take off UPLC Ethernet cable from Quattromicro

- Connect HPLC Ethernet cable to PC Link of Quattromicro

- Power on all machines

- Turn on HPLC and computers of HPLC

- Reboot the Quattromicro

- Windows “start” → go to “run”

- Type “telnet epc”: type “reboot”. If it is ok, it will be a beep signal

- Type “ping epc –t: communication request time out time< 1ms TTL=64

- If that message would be filled in screen, it is Ok.

- Connect the eluent lines

- Take off UPLC eluent line from Quattromicro

- Connect the UPLC eluent line to inlet line of black tube PDA

- Connect the HPLC eluent line to Quattromicro from the outlet line of PDA

- Run the MassLynx

Connect Acquity(UPLC) to Quattromicro

- Power off all machines

- Turn off UPLC, HPLC and computers except Quattromicro

- Connect the UPLC cable

- Take off HPLC Ethernet cable from Quattromicro

- Connect UPLC Ethernet cable to PC Link of Quattromicro

- Power on all machines

- Turn on UPLC and computers of UPLC

- Reboot the Quattromicro

- Windows “start” → go to “run”

- Type “telnet epc”: type “reboot” If it is ok, it will be a beep signal

- Type “ping epc –t: communication request time out time< 1ms TTL=64

- If that message would be filled in screen, it is Ok.

- Connect the eluent lines

- Take off HPLC eluent line from Quattromicro

- Connect HPLC eluent line to the drain line

- Take off UPLC eluent line from inlet line of black tube PDA

- Connect to UPLC eluent line to Quattromicro

- Use the extension line

- Run the MassLynx

- If Quattromicro doesn’t want to use, make a dummy file of tune

- Dummy file of MS tune

- all of voltages is 0

- Source temp is 80 oC and desolvation temp is 30 oC

- Gas flow and Operate are off

- Save the file by dummy file